Purpose

To provide guidelines for continuous cardiac and pulse oximetry monitoring.

Definition

Reference value – The vital sign or hemodynamic measurement used to determine alarm settings, normally the value displayed on the monitor at the time alarms are reviewed and/or set.

Alarm parameter assessment – The comparison of ECG alarm settings (vital signs and/or hemodynamic measures) to patient’s actual vital signs and/or hemodynamic measures to determine if alarm settings should be adjusted.

Policy

- Nurses caring for patients in telemetry-monitored areas will maintain valid AHA BLS and ACLS cards.

- Nurses caring for patients in telemetry-monitored areas will pass an initial and annual ECG exam.

- All patients in intensive care units will be connected to continuous ECG monitoring unless a provider orders otherwise.

- In telemetry, all patients with a provider order will be connected to continuous ECG monitoring or whenever an ACLS trained RN determines it would be useful in an urgent or emergent situation.

- Telemetry box will routinely be set to telemetry mode. In select situations monitor mode may be used (e.g. administration of cardiac medications, change in patient condition)

- When a telemetry box is being used, a safety check must be performed at the beginning of the shift.  The safety check includes the following actions:

  - Check the telemetry box battery life.  Change if battery life is at 20% or less.  NOTE: For patients on continuous pulse oximetry: Check battery life every 4 hours.

  - Switch tele box to monitor mode (if not already in monitor mode) and view the display.  Ensure patient name is visible in the upper left corner.

  - Check alarm limits on tele box or at nursing station.  Alarm limits must be adjusted at the nursing station by the RN.

  - Ensure electrodes are placed correctly. Do not use alcohol to prep the skin for electrode placement.

- All cardiac-monitoring alarms pertinent to the patient's status will be always turned on.  Lead selection is based on reason for monitoring and the patient’s clinical condition.

- Leads III and V2 or V3 are the preferred primary monitoring leads.

  - Patients undergoing revascularization procedures may be monitored in a lead that reflects the target vessel distribution – LAD (V1-4) Circ (V5-6, I, AVL), RCA (II, III, AVF).

- In telemetry, at least one ECG wave should be displayed in the central station. When SpO2 is ordered, the SpO2 should be the second wave displayed. If only SpO2 is ordered, the ECG does not need to be monitored/displayed.

- Lead placement will be verified during the initial assessment on each shift and correlated with ECG morphology on the monitor.

- Primary nurse will ensure telemetry strips are sent to the medical record.

- Patients receiving continuous ECG monitoring must be monitored via transport monitor with defibrillator capability, defibrillator pads and an ACLS Certified RN for the following situations:

  - Off the unit for test or procedure.

  - Transfer to another monitored bed.

  - Away from patient care areas where crash cart with defibrillators is located.

- Monitor Tech (MT) will provide hand off communication during shift change and temporary relief of responsibility.

- At the start of the shift, incoming Monitor Tech will round to introduce themselves verify room number, patient’s name, and telemetry box number and document information on Telemetry Log Shift Handoff Tool and reconciled at central monitor station.

- On admission, RN connects patient to appropriate continuous ECG monitoring equipment and syncs the device to central station. PCT as delegated may assist with telemetry leads placement with telemetry equipment and syncs to central monitor station.

- Only an RN with a valid AHA ACLS card may set and/or adjust alarm limits.

- Upon discharge or discontinuation of ECG monitoring order, RN or PCT removes telemetry equipment (Including ECG leads/electrodes) and communicates with MT that patient taken off telemetry.

- The Primary RN will verbally communicate to the Monitor Tech when a patient is disconnected from telemetry or leaves the floor and when patient is reconnected to telemetry or returns to the floor. RN will verify number on telemetry equipment in patient room matches number displayed on central monitor station.

- Patients may ambulate on IntelliVue X3 monitor on a nursing unit with a code cart//defibrillator present.  Ambulating between immediately adjacent nursing units is allowed (i.e.: 5 East and 5 West, 5 ICU and 5 South). When ambulating off unit or between non- adjacent units, the patient must be monitored on a transport monitor with defibrillator function. (as defined in the Intrafacility Patient Transport policy).

- Transport Monitors must be plugged in charging when not in use.

- Use approved hospital disinfectant wipes to clean MX40 Telemetry boxes.  After contact wet time is achieved, screen residue removal can be performed if needed by wiping with a lint-free damp cloth.  If the inside of patient cable housing must be cleaned, use a small amount of isopropyl alcohol only.

Indications for Cardiac Monitoring

- Class I indications for cardiac monitoring: all patients at significant risk of an immediate, life-threatening arrhythmia. Cardiac monitoring is indicated for most, if not all, patients in this group.

  - Cardiac arrest, early phases of acute coronary syndrome (ACS), unstable ACS/newly diagnosed high risk coronary lesions, acute heart failure, pulmonary edema, AV block, arrhythmias complicating Wolff-Parkinson-White syndrome, long QT syndrome, ventricular arrhythmias, hemodynamically unstable arrhythmias, procedures requiring conscious sedation or anesthesia, cardiac surgery, non-urgent PCI with complications, implantation of automatic pacemaker/defibrillator lead, pacemaker dependent, temporary pacemaker, intra-aortic balloon pump, intensive care indications.

- Class II indications for cardiac monitoring; cardiac monitoring may be of benefit in some patients but is not considered essential for all. Cardiac monitoring often takes place in intermediate care (telemetry) unit.

  - Post-acute myocardial infarction, chest pain syndromes, uncomplicated, non-urgent PCI; subacute heart failure; syncope evaluation. Arrhythmias that cause discomfort in patients with do-not-resuscitate orders. Antiarrhythmic drug or dose adjustment for rate control with chronic atrial tachyarrhythmias. Implantation of pacemaker lead or not pacemaker dependent. Uncomplicated ablation, routine coronary angiography.

- Class III Indications for cardiac monitoring. Cardiac monitoring is helpful in clinical management, but it is not expected to save lives. Cardiac monitoring is not indicated because a patient’s risk of a serious event is so low that monitoring has no therapeutic benefit.

  - Obstetric patients without heart disease. Permanent rate-controlled atrial fibrillation, stable patients with chronic ventricular beats, post-operative patients at low risk for arrhythmias, hemodialysis (except for patients with other Class I or II indications, or those undergoing hemodialysis in the hospital setting.

Monitoring with All Systems

- Ensure that the monitor cable and lead wires are intact with no exposed wiring.

- To admit the patient to the system:

  - Click “Manage Patient”

  - Click “Find Patient”

  - Enter the patients account number (FIN) and click “Search”

  - Select and admit patient

  - For telemetry box and MX40 bedside monitor, assign equipment by clicking on “Equipment” and selecting the appropriate equipment.

- Turn the alarms on and set the upper and lower heart rate alarm limits as required in the “SETTINGS FOR ALARMS” section within this policy.

- Ensure that the monitor senses each QRS complex, increase the gain as necessary.  Ensure that QRS complex is predominantly negative or positive to avoid false alarms.

- Ensure that pacing status is turned on or off as appropriate.

Monitoring with Hard Wire System (bedside, wall-mounted system)

Equipment

  - Cardiac monitor

  - Lead wires, monitor cable

Procedure

- Turn on the monitor while preparing the patient and place in "Standby Mode" where available.

- Attach the lead wires to the electrodes.  See ELECTRODE PLACEMENT AND CARE.  (For “snap on” lead wires, attach to electrodes before placing electrodes on patient.)

  - Set the bedside monitor on "monitor" for routine observation.

  - Use the "filter" mode if it is necessary to filter out artifact that cannot be removed by changing the ECG patches.

  - Use the "diagnostic" mode to enhance waveforms to diagnose dysrhythmia.

Monitoring with The Mx40Telemetry System

Equipment

  - Telemetry box

  - 3 AA batteries

  - 5 lead wire cable set

  - Pulse oximetry cable as needed

  - Telemetry pouch (optional)

Procedure

  - Educate patient regarding rationale for monitoring and how it works.

  - Insert 3 AA batteries in the telemetry box matching the polarity markings on the telemetry box with those on the battery. Close battery compartment and wait for it to automatically turn on.

  - Ensure that the lead wires are intact with no exposed wiring.

  - Connect lead set to telemetry box matching the appropriate terminals to telemetry box.

  - Place electrodes on chest (see electrode placement and care) and attach lead wires.

  - When pulse oximetry monitoring is required in addition to ECG monitoring:

    - Attach the pulse oximetry cable to the SpO2 connection on the telemetry box.

    - Tap SpO2 on telemetry box screen.

    - Select type of monitoring: manual (for spot checks) or continuous.

    - Tap on SpO2 on screen again and scroll down to "Start SpO2".

NOTE: When ECG monitoring only is required, tap on SpO2 on screen and scroll down to “SpO2 on.”     Select/Click to turn off.  When SpO2 monitoring only is required, tap on ECG on screen and scroll down to “ECG on.” Select/Click to turn off.

  - Attach Telemetry box to gown or place in Telemetry pouch.

  - Admit patient (see “Monitoring with All Systems” above.

  - Click “Equipment” and select assigned equipment. NOTE: When an MX40 bedside monitor is available and will be used, ensure both the telemetry box and monitor are selected.

  - Verify rhythm and heart rate on central station match patient’s expected rhythm and heart rate.

  - Verify number on telemetry equipment in patient room matches number displayed on central station.

Monitoring with Philips Intellivue x3

Equipment

- IntelliVue X3 monitor

- Dual Adaptor

- Claw mounts

Procedure

            - Explain to patient rationale for monitoring during ambulation and how it works.

- Ensure that the lead wires are intact with no exposed wiring.

- Ensure electrodes are attached on chest.

- Attach monitor to pole.

- Turn monitor on using the front right button.

- Remove ECG cable, O2 cable, 1-2 pressure cable (if applicable) from bedside monitor and attach to portable monitor.

Note: if monitoring two pressure lines, use dual adaptor.

- Using the touch screen, press screen key to enable touch screen use.

- Using touch screen adjust alarms, vital trend, patient demographic and profiles as needed, Adjust pressure waveform scales and settings by pressing keys then main set up, then measurements.  Select the pressure monitoring parameter you wish to adjust.  Touching the screen key brings you back to the display main screen.

- After ambulation, detach cables from IntelliVue X3, and reattach to bedside monitor, ON IntelliVue X3 press Keys then end case to erase patient data. Press power key to turn off and return to charge.

Transfer/Discharge (all system)

  - If the patient is transferring to a monitored unit, arrange for transport with continuous ECG monitoring and defibrillator capability.

  - Discharge patient from the monitor upon transfer/discharge. The monitor automatically saves data with discharge.

Monitor Tech Alarm response

    - When patient becomes disconnected (even partially), from the telemetry monitor system, develops dysrhythmias, or has a questionable rhythm the MT will:

    - Notify Patient Care Technician (PCT) when a patient needs to be connected or have the telemetry monitoring system checked.

    - If the PCT is unavailable, the MT will notify the primary RN or the resource RN.

    - If the PCT, primary RN, and resource RN are busy, the monitor tech will notify the charge RN.

    - If all the above are busy, any available RN is notified to immediately check on the patient.

    - Once patient is reconnected, the PCT/RN must call MT to verify patient is appropriately connected.

    - In any case where dysrhythmia is identified, the MT will need to notify the RN immediately.

Electrode Placement and Care

Equipment

  - Disposable pre-gelled electrodes

  - Shaving equipment

  - Gauze, dry wash cloth or exfoliating tape

Procedure

  - Electrode locations

    - 5 lead cable system

      - RA (white) – Right mid-clavicular line, just below clavicle

      - LA (black) – Left mid-clavicular line, below clavicle

      - RL (green) – Right anterior axillary line, below the level of the diaphragm

      - LL (red) – Left anterior axillary line, below the level of the diaphragm

      - V (brown) – 4th intercostal space, left sternal border (V2), or 4th intercostal space, right sternal border (V1). V2 is the preferred default for ST monitoring.

Standard 5 lead with V2 configuration

    - Standard 3-Lead, Lead II

        - RA (white) – Right mid-clavicular line, below clavicle

        - LA (black) – Left mid-clavicular line, below clavicle

        - LL (red) – below the diaphragm, near anterior-axillary line

Standard 3 lead, Lead II placement

    - MCL1 (MCL=Modified Chest Lead)

        - RA (white) – Left mid-clavicular line, below clavicle

        - RL (red) – 4th intercostal space, right sternal border

        - LA (black) – Right mid-clavicular line, below clavicle

    - MCL6

        - RA (white) – Left mid-clavicular line, below clavicle

        - RL (red) – 5th intercostal space, Left mid-clavicular line, mid axillary line

        - LA (black) – Right mid-clavicular line, below clavicle

3 lead, MCL1 & MCL6 configuration

  - Apply electrodes

    - Clean the intended the electrode areas with soap and water.  Then dry completely with dry gauze or washcloth.

    - If needed, clip or trim the chest hair where the electrodes are to be placed.

    - Exfoliate the skin using gauze, dry washcloth, or exfoliating tape.

    - If the patient is diaphoretic, apply Benzoin and allow to dry.

    - Apply the electrodes firmly to the appropriate locations

  - Electrode Care

    - Assess skin integrity around the electrodes every shift.

    - Monitor the patient’s skin for signs of irritation or allergic reaction to electrodes.  Change electrodes daily if any irritation occurs and keep sites as clean and dry as possible.

    - Change the ECG patches when a signal problem is evident, or when patches become soiled.  It is recommended that all electrodes be changed at the same time.

Reviewing and Saving Telemetry Strips

            - Strips will be saved

              - Every 12 hours (0800 and 2000)

              - With changes in rhythm or condition

              - For significant events (i.e., red alarms, Code Blue).

            - On units with monitor techs, the monitor tech will save the strips and annotate the rhythms and measurements within the Philips monitoring system.

              - The RN must send the strip to EMR after reviewing and verifying the information is correct

              - The RN will then log into KeckCare and following the instructions below to acknowledge the strip.

            - Reviewing and saving alarms:

              - In patient sector, click on REVIEW- This brings you to alarm review.

              - If you wish to send an "Alarm Strip" to EMR, highlight the alarm (click on the alarm you want)

              - To annotate select "Calipers" at bottom of Strip window or click on “Annotate Icon” measure/annotate as desired- SAVE

            - Saving routine strips:

              - Go to Alarm Review.

              - At bottom of screen, click on REVIEW and you'll have options for different reviews.

              - Go to General Review

              - Select the wave that you would like (hovering will show you the exact time of the strip, or you can select a time by clicking on the tabular column or timeline at bottom of screen). -Click and it will open the "Strip Window"

              - Once the strip is open, complete your measurements:

              - Click on "Calipers" at bottom of Strip window or click on “Annotate Icon”

              - Measure as needed.

              - Click "Save" and it will be saved for the time the strip occurred, under "Alarm Review"- Indicated by Yellow Text Bubble

            - To Export this to EMR, click on the small printer icon on the upper right of the strip window.

            - From the drop down select Electronic Document (If you want a printout select Paper).

              - NOTE: Monitor techs may save, annotate, and measure strips, monitor techs may not send strips to Cerner.

              - The RN must send the strip to EMR after reviewing and verifying the information is correct.

            - Reviewing and acknowledging saved strips from EMR:

              - Login to EMR then open the patient chart

              - Go to Reports/Documents section of the EMR

              - The saved strips would be saved as telemetry and rhythm Strips document type under Cardiovascular Diagnostics/Noninvasive Cardiology

              - The RN will acknowledge the strip in EMR. With the document open, right-click on the blue PDF link and click “Review” this will add the RNs name to the document showing that the RN reviewed the strip.

            - If the electronic strip sending system is down, continue to save strips and send to EMR when system is back online.

            - If the patient is discharged prior to system coming back online, print strips and place in hard chart.

              - To retrieve saved strips to send to EMR:

                - Go to “Patient Window”.

                - Click “Review”, then “Alarm Review”

                - At the top right, just under the red “X” button, click “Change Tile” button (small button with two opposite facing arrows), and change to “Compressed”.

                - Click “Filter Alerts” (bell button under “Change Tile”) and choose “Saved Strips”.

                - Click "Timeline Set-Up” (bottom right) to change timeline duration.

                - Scroll back and forth using arrows at the bottom of the screen to search for dates/times that need to be sent to EMR.

ST-Segment Monitoring

- ST segment monitoring is defaulted to “ON”.

- The bedside monitor will automatically detect and set an ST segment reference value.

- Alarm settings:

    - In critical care: 1mm above/below patient reference value.

    - In PACU, procedural areas and on telemetry floors: 2 mm above/below patient reference value.

    - Telemetry floor central stations do not have the ability to default to only two leads, thus all ST segments are monitored. Lead III and V are defaulted to 2mm above/below the reference value. All other leads are defaulted to 4 above and below reference value and may be set in addition to leads III and V based on RNs clinical judgement.

- ST-segment monitoring can detect silent ischemia, which occurs in the absence of symptoms.  ST segments should be turned off for patients with:

    - Conduction disturbances such as: Left Bundle Branch Block, Wolff-Parkinson-White Syndrome, Intermittent Right Bundle Branch Block.

    - Ventricular pacing.

    - Atrial fibrillation with coarse isoelectric line.

    - Resting ST segments greater than 1mm- especially depression.

    - Frequently agitated or restless patients.

    - Per provider order.

- Multiple non-ischemic causes of ECG changes can mimic ischemia through ST elevation or depression, such as:

    - Body position changes

    - Drug effects (such as digitalis and diuretics)

    - Electrolyte imbalances (such as hypokalemia)

    - Conduction disturbances (including LBBB and WPW syndrome)

    - Hypothermia

    - Left ventricular hypertrophy

    - Ventricular pacing

- ST segment reference value should NOT be routinely reset- the goal is to trend ST segment elevation or depression.

- You should reset the reference value if:

    - The ECG leads were positioned incorrectly when the reference value was set

    - the patient was agitated or moving when the reference value was set

    - the patient had a new/evolving MI while in the hospital

    - Per provider order

    - If your ST segments have questions marks next to them, you may need to increase the gain (size up) of your ECG leads. It is difficult for the monitor to detect the ST segment in patients who have low-voltage ECGs.

Note: Because ST segments may be altered by several non-ischemic conditions, it is imperative that clinicians interpret the ST information in the context of other clinical information.

- Response to ST segment alarms:

      - Verify that alarms are correctly set for one mm above/below the isoelectric line for critical care and 2mm above/below the isoelectric line on telemetry units.

      - Place patient in a supine position without excessive movement.

      - If alarm condition continues notify provider and obtain an order for a 12 lead ECG to confirm ST changes. ST depression or elevation of 1 to 2 mm that lasts for at least 1 minute can be clinically significant and warrants further patient assessment.

QT Monitoring

            - QT monitoring is defaulted to ON

- Some patients are at increased risk for Torsade de Pointes. Females, age greater than 65, bradycardia, impaired left ventricular function (ischemia, left ventricular hypertrophy), hypokalemia, hypocalcemia and hypomagnesemia are all conditions which increase the risk of Torsades de Points. A prolonged QT interval may be a warning sign.

- Measurement Definitions

- QT:  QT interval in milliseconds. The QT interval is defined as the time between the beginning of the Q-wave and the end of the T-wave.

- QTc:  The heart rate corrected QT interval. Faster heart rates shorten the QT interval and slower heart rates prolong the QT interval.

- ΔQTc:  The delta QTc is the difference between the current QTc value and the QTc reference value.

- QT-HR: The heart rate used to calculate the QTc value

- The following is list of drugs that may prolong the QT interval, there are medications outside of this list that may also prolong QT intervals.

    - Antiarrhythmic drugs: Quinidine, Procainamide, Disopromide, Dofetilide, Ibutilide, Sotalol

    - Antipsychotic and antidepressant agents: Haloperidol, Droperidol, Thioridazine, Chlorpromazine

    - Atypical antipsychotics: Ziprasidone, Risperidone, Zimeldine, Citalopram

    - Antidepressants: Amitriptyline, Desipramine, Imipramine, Maprotiline, Doxepin, Fluoxetin

    - Antibiotics: Levofloxacin, moxifloxacin, Erythromycin, Clarithromycin

    - Antimalarials: Quinine, halofantrine

    - Antiprotozoal: Pentamidine

    - Antifungal: Azole group

    - Methadone

    - QT Measurement Limitations

          - Under the following conditions it may be difficult to achieve reliable QT monitoring and may produce a “Cannot Analyze QT” INOP message. During QT startup period, this INOP is announced immediately when QT value is invalid. After startup, this INOP is announced after the QT measurement is invalid for greater than 10 minutes.

          - T-wave detection limitations: To accurately measure the QT interval, the algorithm must determine the beginning of the QRS and the end of the T-wave. The following conditions will make it difficult to determine the end of the Twave:

                - Flat T-wave

                - Atrial Fibrillation or Atrial Flutter

                - Prominent U-waves

                - In cases where the QT cannot be measured reliably, select a single lead with a good T- wave amplitude and no visible fibrillatory or flutter activity, and without a predominant U-wave or P-wave.

          - QRS Changes

          - A widened QRS duration contribute to a long QTc interval. In this case, if a long QTc is observed you should verify it to ensure that it is not caused by QRS widening.

          - Rhythm and rate limitations

          - High heart rate:

          - When the heart rate is high, the P wave of the subsequent P-QRS-T complex may approach the end of the T wave and produce an effect like that of a U wave.

                - – Heart rate greater than150 beats/min for adult patients

                - – Heart rate greater than 180 beats/min for pediatric or neonatal patients.

          - Paced Rhythm: Ventricular Paced or AV sequentially paced rhythm produced widened QRS

          - Bigeminy Rhythm: May produce a longer QT interval due to widened QRS on every other beat.

          - In these cases if rhythm is sustained you may consider turning QT Interval monitoring off.

          - Low voltage ECG- try increasing the gain (size up) of the ECG leads.

          - Response to QT alarms

              - Verify that none of the above conditions are present.

              - Review the patient’s medication list for possible causes.

              - Ensure that the QT was set at 20% above/below reference value.

              - Notify provider if alarm condition persists and obtain 12-lead ECG.

Pulse Oximetry Monitoring

- Apply the pulse oximeter probe to the patient’s finger and connect to the transmitter if indicated.

- Avoid IV arm or extremity w/edema if possible.

- Preferred placement of digit sensor is the middle or ring finger of the non-dominant hand, or the large or second toe.

- If patient is gripping, as with the use of a trapeze use ring finger, pinky, or toe.

- Avoid digits with dark or metallic nail polish.

- The following patients should receive continuous pulse oximetry monitoring via provider order:

  - All ICU and stepdown patients

  - Patients on oxygen therapy greater than 4L nasal cannula- all other forms of oxygen therapy including CPAP, BiPAP, simple mask, etc. require continuous pulse oximetry

  - Patients with a tracheostomy

  - Patient-controlled analgesia or epidural therapy

  - First 24 hours post-operatively for patient’s w/history of cardiac or pulmonary disease

  - First 24 hours post-operatively for patients undergoing abdominal or thoracic surgical procedures

Telemetry Alarms

- Level 1 Critical RED Alarms:

  - Asystole.

  - Ventricular Fibrillation.

  - Ventricular Tachycardia.

  - 3rd Degree Heart Block.

  - Failure to Capture (alarms as yellow but should be regarded as a critical “RED” alarm).

  - Failure to Pace (Alarms as yellow, but she be regarded as a critical “RED” alarm).

  - Apnea.

  - Desaturation less than 90 %.

  - Any other dysrhythmia that alarms RED on the cardiac monitor.

  - Alarms for life-threatening rhythms (RED alarms) will be always on.

Upon notification of red alarms the Monitor tech will:

  - Immediately alert any nursing or medical staff in the immediate vicinity to respond and evaluate the patient.

  - Notify the primary RN.

  - Once notified, Monitor tech documents the name of RN onto strip and saves in central station.

  - Primary RN will verify and review strip and send to EMR.

  - If no immediate response from primary RN, Monitor tech will escalate to nurse lead.

- Level 2 Yellow Alarms:

  - Technical YELLOW alarms include all other alarms that are not included in Level 1.

Upon notification of yellow alarms the Monitor tech will:

  - Send a text page to primary RN.

  - If no resolution after five minutes, monitor tech will send a text page to resource RN/NL.

- Alarms should be temporarily suspended when the RN is at the bedside, directly visualizing the patient and performing a procedure that may cause false alarms. (i.e., removing ECG leads for a bath, drawing a blood sample from an arterial line, etc.)

Settings for Alarms

- To reduce the number of nuisance/non-actionable alarms, all telemetry alarm parameters will be individualized to the patient.

- Alarm parameters will be assessed and set upon admission and/or initial connection to ECG monitoring.

- The following alarm parameters will be set to no more than 20% above and below the reference value.

  - Heart rate

  - Blood pressure

  - Threshold for “PVC Rate”

  - QT intervals

  - Threshold for “PVC”

- An alarm parameter assessment will be performed each shift, and prn (e.g., change in patient condition or when alarm parameters are exceeded/subceeded).

- Alarm parameters will be adjusted when the alarm parameter assessment indicates the vital sign or hemodynamic value displayed on the monitor have changed since the last alarm parameter assessment (e.g. value now outside of set alarm parameters).

- Alarm parameters that should be individualized by turning them on or off include but are not limited to:

  - Non-sustained VT

  - Pair PVCs

  - Ventricular bigeminy/trigeminy

  - Atrial fibrillation

  - Irregular heart rate

  - ST segment monitoring

Documentation

Rhythm will be documented in EMR with each set of vital signs.

See “reviewing and saving telemetry strips” above.

Nursing Response to Equipment Failure

- Notify Biomedical Engineering of the problem.

  - Use an extra transmitter or wireless monitor when available.  If an extra transmitter is unavailable, and patient requires continuous ECG monitoring, the patient will be transferred to an available monitored bed at appropriate level of care.

  - Monitoring equipment in ICUs – see Safety Manual

- If the electronic strip saving system is down, continue to save strips and send to EMR when system is back online.

- If the patient is discharged prior to system coming back online, print strips and place in hard chart.

  - To retrieve saved strips to send to EMR

      - Go to “Patient Window”

      - Click “Review”, then “Alarm Review”

Reference(s)

- AACN, (2016). AACN Practice Alert- ST Segment Monitoring.

- Carlson, K. and Lynn-McHale Wiegand, D. AACN Procedure Manual for Critical Care (5th edition). Elsevier Saunders 2005.

- Crawford, C.L. (2015). Telemetry Monitoring: Are admission criteria based on evidence? American Journal of Critical Care, 24(4), pg. 360-364.

- Lippincott Procedures - Cardiac monitoring (lww.com)

- Phillips Corporation (2013). Just a Nuisance.

- Phillips Corporation (2012). ST Segment Monitoring: ST/AR Arrhythmia Algorithm Application Note

- Drew, BJ., Harris, P., Zegre-Hemsey, JK., et. al. (2014). Insights into the Problem of Alarm Fatigue with Physiologic Monitor Devices: A Comprehensive Observational Study of Consecutive Intensive Care Unit Patients. PLoS One, 9(10).

- Drew, BJ., Califf, RM., Funk, M. et al. (2004). AHA Scientific Statement: Practice standards for electrocardiographic monitoring in hospital settings. Circulation, 110:2721-2746.

- Nachimuthu, Assar, M., Schussler, J. (2012). Drug-induced QT interval prolongation: mechanisms and clinical management. Therapeutic Enhancements in Drug Safety, 3(5): 241–253.

Attachment(s)

- Telemetry Log (0700-1900) Shift Handoff Tool

- Telemetry Log (1900-0700) Shift Handoff Tool

| AACN’s Evidence-leveling System. Level A; Meta-analysis of multiple controlled studies or meta-synthesis of qualitative studies with results that consistently support a specific action, intervention or treatment. Level B; Well designed controlled studies, both randomized and nonrandomized, with results that consistently support a specific action, intervention, or treatment. Level C; Quality studies, descriptive or correlational studies, integrative reviews, systematic reviews, or randomized controlled trials with inconsistent results. Level D; Peer-reviewed professional organizational standards, with clinical studies to support recommendations. Level E; Theory-based evidence from expert opinion or multiple case reports. Level M; Manufactures’ recommendations only. | ##### Effective/Revision Dates for Policy # NA 10-129A. Effective:; 05/02/1991; MEC. Revised:; 03/12/1996; Renumbered. Revised:; 04/14/1998; MEC. Revised:; 05/03/2002; Professional Practice. Revised:; 03/29/2005; Governing Board. Revised:; 09/30/2010; Governing Board. Revised:; 03/05/2015; Governing Board. Revised:; 10/27/2016; Governing Board. Revised:; 11/30/2017; Governing Board. Revised:; 02/28/2019; Governing Board. Reviewed:; 03/22/2022; Policy Committee. Revised:; 12/29/2022; Quality Committee of the Governing Board. Revised:; 10/17/2024; Quality Committee of the Governing Board. Revised:; 05/27/2025; Governing Board. 00/00/0000. Keywords:; Telemetry, ECG, Monitor. |
| --- | --- |