Purpose

The purpose of this policy is to outline the organizations standard for temporary pacing using epicardial wires.

Definitions

  - Milliamps (mA): The electrical output of an artificial pacemaker causing depolarization of the atria and/or ventricle.

  - Sensitivity: The degree or amplitude the pacemaker "sees" or senses intrinsic heart activity. Sensitivity is measured in millivolts (mV).

  - Rate: In fixed or asynchronous mode, the pacemaker will pace at a set rate regardless of the patient’s intrinsic rhythm. Asynchronous pacing should only be utilized in the absence of an intrinsic rhythm. In demand mode, the pacemaker will pace only if the patient's intrinsic rate falls below the set rate.

Policy

- An extra pulse generator will be readily available on the unit.

- Epicardial wires will be insulated (i.e., placed inside a glove) and dressed at all times.

- The epicardial wires will not be disconnected from the pulse generator without a provider order. When the generator is disconnected, the bridging cable (adapter) will remain connected to the epicardial wires and will be insulated, dressed, and taped to the patient’s abdomen/chest for emergent access.

- Dressing will be changed every 24 hours and prn.

- A new battery will be placed in the pulse generator prior to initiating use.

- Circumstances under which emergency institution may be performed (see Standardized Procedure for Emergency Care and Arrhythmia Treatment): Electrical pacing shall be instituted by a nurse for patients with temporary epicardial wires when the patient's heart rate slows to less than 50 BPM and bradycardia is symptomatic as evidenced by one or more of the following signs or symptoms:

  - Hypotension

  - Signs and symptoms of cerebral insufficiency

  - Dyspnea

  - Chest pain

  - Occurrence of six or more "rate-related" PVC's per minute

  - Second- or third-degree heart block

- Notify provider when pacing is initiated.

- Nurses who have completed the following may institute temporary epicardial pacing.

  - Advanced Cardiac Life Support (ACLS) trained

  - Emergency Pacing Skill validation

Continuing Care

- Monitor the mA setting

  - Insufficient mA may cause failure to capture.

  - Excessive mA may cause irritability, leading to lethal dysrhythmia and can shorten the life of epicardial wires.

  - Note: Atrial and ventricular thresholds for epicardial wires often increase by the fourth post-operative day.

- Monitor the Sensitivity Setting

  - Excess sensitivity may inappropriately inhibit the pacemaker due to sensing of extra-myocardial activity (i.e., muscle movement, artifact) as myocardial activity, resulting in a failure to pace.

  - Turning the sensitivity down too low (higher number= lower sensitivity) may cause fixed or asynchronous rate pacing with the possibility of the pacer stimulus being delivered during the vulnerable period of the ECG cycle (i.e., the T wave) leading to lethal dysrhythmia.

  - Asynchronous or fixed pacing should only be utilized in the absence of an intrinsic rhythm.

- Monitor the Rate

  - Excessive rate may cause competition with patient’s own intrinsic rate, again possibly leading to lethal dysrhythmia.

  - Insufficient rate may lead to insufficient cardiac output. Normal rate = 60 - 100 BPM.

- Monitor the patient

  - Monitor vital signs and hemodynamic response to pacing per provider order

  - Evaluate ECG for presence of paced rhythm or resolution of initiating dysrhythmia after patient is initiated.

  - Notify provider of any pacemaker problem and corrective actions taken.

- Use Safety Measures: Do not touch the pacing wires with wet hands, and do not allow fluids (such as H2O, IV fluids, emesis, etc.), to come in contact with the pacing system. Ensure the ends of the wires are secured into the bridging cable, and that the cable is secured into the terminals of the pulse generator at all times.

  - Keep bridging cable/epicardial wire connection insulated in a rubber glove.

  - Take precautions to ensure the pulse generator will not fall or otherwise pull on the pacing wires or become disconnected.

  - Instruct the patient not to use an electric razor or any other nonessential electrical equipment.

  - Wear rubber gloves when handling the pacing wires to prevent microshocks.

  - Insulate exposed metal terminals on a pacemaker with rubber sleeves.

INSTITUTING PACING

Equipment

- Temporary external pacemaker with fresh battery

- Disposable bridging cable (also known as an “adapter”)

- Gloves

- Continuous electrocardiogram (ECG) monitor

- 12 Lead ECG electrode (as needed for unipolar pacing)

- Alligator clips (as needed for unipolar pacing)

Procedure

- Ensure patient is connected to continuous ECG monitor.

- Bipolar Electrodes and Epicardial Wires (both leads are in contact with the myocardium and able to function as negative poles).

  - Put on clean gloves to protect patient from microshock.

  - Insert new battery into pulse generator and check for battery signal. Record on battery the date/time of new battery insertion.

  - Insert atrial wires to (+) and (-) terminal of atrial pacing bridging cable. Insert ventricular wires to + and - terminal of ventricular pacing bridging cable.

  - Connect the bridging cable to the pacemaker pulse generator. For a dual chamber pacemaker, ensure the atrial wires are inserted on the side marked “A”, and the ventricular wires are inserted on the side marked “V”.

  - Turn generator "on".  Observe the ECG monitor.

  - Place pacemaker generator in ordered mode and settings ordered by the provider. If initiating pacing and no orders are present, start at the lowest settings on the pulse generator.

  - If 100% capture not present, increase the mA level until 100% capture is noted. Capture is indicated by an atrial and/or ventricular response with each pacemaker impulse. To prevent fibrosis, the lowest safe mA setting should be utilized. Check pulse to ensure mechanical capture.

  - Determining the sensitivity threshold (level at which intrinsic myocardial activity is recognized by the sensing electrodes). Perform this test with the pacer rate set at ten below the patient’s intrinsic rate and the mA at the lowest possible setting:

    - Do not perform a sensitivity threshold test if patient has no intrinsic rhythm.

    - Turn sensitivity down (counterclockwise, to a higher number) until sense indicator light stops flashing. This is when the pulse generator is no longer able to detect the patient’s intrinsic rhythm.

    - Slowly turn the sensitivity back up (clockwise, to a lower numeric setting) until the sense indicator light flashes with each complex and the pace indicator light stops- this is the sensing threshold.

    - Set the sensitivity dial to the number that was half the sensing threshold.

  - Determining the stimulation threshold (the amount of output required to achieve capture):

    - This test should be performed by a provider if the patient is pacemaker dependent for bradycardia.

    - Set the pacing rate 10 beats above the patient’s intrinsic rate.

    - Gradually decrease the mA from 20 until capture is lost.

    - Gradually increase mA until 1:1 capture is established. This is the stimulation or pacing threshold.

    - Set the mA at least two times above the stimulation threshold. For example, if 1:1 capture is achieved at three mA, then mA on pulse generator should be set at six mA.

  - Monitor vital signs and patient symptoms until vital signs are stable and symptoms have resolved.

  - Isolate exposed ends of pacing wires by placing a rubber glove over the cable-wire connection.

  - Notify provider of initiation of electrical pacing and the patient's status.

  - Obtain new provider orders if settings necessary for pacing are different from original order (i.e., higher mA is needed for capture than originally ordered).

- Unipolar Electrodes (one lead is in contact with the myocardium and the other lead is in placed on the chest wall- this is the ground or skin lead.)

  - Put on clean gloves to protect patient from microshock.

  - The single epicardial wire is attached to the negative terminal of the pacing wire cable.

  - A ground wire, which has been placed subcutaneously, is connected to the positive terminal of the pacing wire cable.

  - In the case that there is no wire in place to be used as a ground (second ventricular epicardial wire has been pulled out), follow the procedure below for establishing a temporary ground electrode:

    - Attach 12 lead ECG electrodes to the patient's chest.

    - Attach one end of an alligator clamp to the chest electrode.

    - Attach the other end to the positive pole of the pacing cable.

    - Notify the appropriate provider immediately to place a wire suture in the skin to be used as the grounding electrode, or to place a new pacing catheter.

    - If pacing cannot be instituted with an ECG electrode and alligator clamp, a subcutaneous needle can be inserted into the chest wall near the apex of the heart and taped firmly in place with enough metal exposed so that an alligator clamp can be attached. Proceed with noninvasive pacing if unable to establish a grounding electrode. See Pacing: Noninvasive External Policy

  - Continue as above in procedure B. 3 through 11.

- Biventricular Pacing

  - Some patients come out of the OR (heart surgery) with biventricular pacing wires or three ventricular wires. They come connected to an adapter so that there are still two leads to the pacemaker cable, but otherwise functions the same as regular ventricular wires. NOTE: Please do not discard adapter.

DRESSING CHANGE

Equipment

- Sterile 4 x 4 dressing gauze

- Tape

- 2% chlorhexidine sponge

- Sterile gloves

- Non-sterile gloves

Procedure

  - Explain procedure to patient

  - Perform hand hygiene, apply gloves

  - Remove old dressing and discard dressing and gloves

  - Perform hand hygiene, apply sterile gloves

  - Observe epicardial exit site for signs and symptoms of infection

  - Apply 2% chlorhexidine sponge in a back and forth scrubbing motion on the surrounding skin starting at exit site and extending outwards, being careful not to place friction over wire that could dislodge it. Allow to dry for 1 minute.

  - Apply sterile 4 x 4 dressing gauze and tape to secure

  - Dressing will be changed every 24 hours, or immediately if the dressing is damp, wet, or no longer intact.

DISCONNECTING WIRES FROM GENERATOR

Equipment

- Non-sterile gloves

- 4 x 4 dressing gauze for each set of wires

- Paper tape

- Plastic tape

Procedure

  - If insertion site dressing change needed, follow steps above for dressing change.

  - Perform hand hygiene, apply gloves.

  - Unplug disposable extension cable from pacing generator, leave epicardial wires plugged into black plastic adapter.

  - Gently roll wires and adapter up and place isolated wires on top of 4 x 4 dressing gauze covering insertion site. Cover with second 4 x 4 dressing gauze and secure with paper tape.

  - Label dressing with type of wires (atrial or ventricular), date, time and initials.

Removal

Removal of epicardial wires will be performed by a provider:

  - Wires must be pulled with the patient in the bed in a supine position.

  - Bedrest will be maintained for one-hour post-removal.

  - Monitor the patient for signs and symptoms of cardiac tamponade after wire removal.2

  - Vital signs will be monitored post-removal:

    - Q 15 minutes x 4

    - Q 30 minutes x 2

    - Q one hour x 2

    - Then resume routine vital signs per provider order.

    - Disconnecting from the pacing generator and wire removal will be discussed daily during rounds.

    - Whenever possible, wires should be removed at least 24 hours prior to discharge.

Documentation

Document the following in the patient’s medical record:

  - Reason for initiation of temporary pacing.

  - The patient's response to the procedure.

  - Mode, mA, rate, and sensitivity setting.

  - On initiation of pacing: ECG strips labeled with date, time, patient's initials, evidence of need for pacing, capture beats, and evidence of appropriate sensing if possible.

  - After initial strip, ave and ECG strips with date, time, pacemaker type and settings per ECG Monitoring policy.

  - Appropriateness of capture and sensing.

  - Provider notification and corrective actions taken if capturing or sensing problems occur.

  - Dressing change time and date.

  - Disconnection from pulse generator.

  - Removal of epicardial wires.

Related Policies and Procedures

- Nursing Manual

  - Standardized Procedure for Emergency Care and Arrhythmia Treatment

  - ECG Monitoring

Appendix

- Appendix A: Pacemaker Malfunctions and Troubleshooting

Reference(s)

- Van Blarcom AG, Wojack CA, Casida J. Cardiac Tamponade Following the Removal of Epicardial Pacing Wires: Critical Care APRN Toolkit. AACN Adv Crit Care. 2020 Dec 15;31(4):410-415. doi: 10.4037/aacnacc2020324. PMID: 33313709.

- Safavi-Naeini P, Saeed M. Pacemaker Troubleshooting: Common Clinical Scenarios. Texas Heart Inst J. 2016 Oct 1;43(5):415-418. doi: 10.14503/THIJ-16-5918. PMID: 27777525; PMCID: PMC5067035.

- Wiegand, Debra. AACN Procedure Manual for Critical Care 7th Edition, pp 417-429. St Louis, MO, Saunders, 2017.

| AACN’s Evidence-leveling System. Level A; Meta-analysis of multiple controlled studies or meta-synthesis of qualitative studies with results that consistently support a specific action, intervention or treatment. Level B; Well designed controlled studies, both randomized and nonrandomized, with results that consistently support a specific action, intervention, or treatment. Level C; Quality studies, descriptive or correlational studies, integrative reviews, systematic reviews, or randomized controlled trials with inconsistent results. Level D; Peer-reviewed professional organizational standards, with clinical studies to support recommendations. Level E; Theory-based evidence from expert opinion or multiple case reports. Level M; Manufactures’ recommendations only. | ##### Effective/Revision Dates for Policy # NA 10-154A. Effective:; 05/02/1991; MEC. Revised:; 04/14/1998; MEC. 12/2001; Critical Care Committee. 04/05/2002. 05/24/2005; Governing Board. 06/05/2007; Governing Board. 01/27/2011; Governing Board. Revised:; 07/12/2016; Policy Committee. Revised:; 12/13/2016; Policy Committee. Revised:; 01/28/2020; Policy Committee. Revised:; 02/14/2023; Policy Committee. Revised:; 03/10/2026; Policy Committee. 00/00/0000. Keywords:; Pacing, Pacemaker, Arrhythmia, epicardial, pacer, wires, cardiac surgery. |
| --- | --- |

APPENDIX A: Pacemaker Malfunctions and Corrective Actions

  - Failure to capture: Generated pacing stimulus does not initiate myocardial depolarization (pacemaker spike with no QRS or P wave after it).

- Possible causes include:

  - Lead dislodgment, low output, lead maturation (old leads), and lead or pacer failure. Epicardial wires have a higher incidence of failure after two weeks.

- Possible solutions:

  - Change patient's ECG lead to see if capture is detected in another lead (12 Lead ECG may be helpful).

  - Increase mA in increments of 1 until 100% capture achieved.

  - If problem continues, change pulse generator.

  - Failure to pace: Heart rate is below the programmed lower rate of the pacemaker and no pacer electrical output is noted on the electrocardiogram.

    - Possible causes include lead failure (lead fracture, lead displacement), pulse generator failure, battery failure, crosstalk inhibition, or oversensing.

    - Possible solutions:

- Replace battery.

- Change pulse generator.

- Double check epicardial wire to bridging cable connection.

- If the above does not resolve issue, the pacemaker may be oversensing, see below for oversensing.

  - Oversensing: Occurs when the pacemaker senses electrical signals that it should not sense, which results in inappropriate inhibition of the pacing stimulus. Oversensing resembles failure to pace/ on the monitor. If adjusting the sensitivity corrects the issue, the pacemaker is likely oversensing.

        - Possible causes: T waves, non-physiologic signals like electromagnetic interference, lead failure.

        - Possible solutions:

          - Ensure sensitivity is not set too high- turn the sensitivity down (note, when turning the sensitivity down, the number (in mV) will go up. Caution: turning sensitivity all the way down (to the highest number) will result in asynchronous pacing.

  - Undersensing: pacemaker fails to detect spontaneous myocardial depolarization, which results in asynchronous pacing. Atrial or ventricular pacing spikes arise regardless of P waves or QRS complex. Results in the appearance of too many pacing spikes or inappropriate pacing spikes (i.e., on the QRS).

  - Possible causes: Pacemaker programming problems (improper sensing threshold), insufficient myocardial voltage signal, lead, or pacer failure (fibrosis, fracture, etc.), or an electrolyte abnormality.

  - Possible solutions:

  - Ensure sensitivity is not set too low (higher number is lower sensitivity) or asynchronous pacing may occur. (The lower the number, the higher the sensitivity). Slowly turn sensitivity up (the number on the pacemaker will go down) until problem resolves. If problem continues, change battery.

  - If problem continues, change pulse generator.