Purpose

The purpose of this policy is to outline the organizations standard for the use and care of temporary transvenous pacing wires.

Policy

- Circumstances under which emergency institution may be performed (see Standardized Procedure for Emergency Care and Arrhythmia Treatment): Electrical pacing shall be instituted by a nurse for patients with temporary epicardial wires when the patient's heart rate slows to less than 50 BPM and bradycardia is symptomatic as evidenced by one or more of the following signs or symptoms:

  - Hypotension

  - Signs and symptoms of cerebral insufficiency

  - Dyspnea

  - Chest pain

  - Occurrence of six or more "rate-related" premature ventricular contractions per minute

  - Second- or third-degree heart block

- Notify provider when pacing is initiated.

- Nurses who have completed the following may institute temporary epicardial pacing.

  - Advanced Cardiac Life Support (ACLS) trained

  - Transvenous Pacing Skill Validation

- Transvenous wires may be removed with a provider order by RNs competency validated in the removal of pulmonary arterial lines. If a permanent pacemaker is in place, the transvenous wire must be removed by a provider.

- An order must be obtained to transfer a patient with a transvenous pacemaker outside of critical care. Only floors where nurses have been trained in the care of the transvenous pacemaker patient may receive patients with a transvenous pacemaker:

  - The patient must have an underlying/intrinsic rhythm: patients who are pacemaker dependent may not be transferred outside of critical care.

  - Transvenous pacemakers inserted in the jugular or subclavian veins are allowed outside of critical care. Pacing wires inserted in the femoral vein must be cared for in a critical care unit.

  - Stopcocks will be removed from the introducer if possible.

  - The introducer and pacing wires must be removed by a competency-validated critical care RN or provider when pacing is no longer needed.

- The sideport of the introducer should have fluids running to maintain patency (See ‘Standardized Procedure for the Ordering of Flushing Vascular Access Device’ policy).

  - Medications may be infused through the sideport.

- Insertion of temporary pacemaker wires may be done in the patient's room in the critical care areas.

- New pacemaker batteries are to be dated and timed placed in the pulse generator prior to each use.

- An extra pulse generator will be readily available on the unit.

Care of Patient/Pacing System

  - Monitor vital signs and hemodynamic response to pacing at least every hour and prn in critical care, every four hours, and prn on telemetry or stepdown.

  - Do not touch the pacing electrode with wet hands, and do not allow fluids such as H2O, IV fluids, emesis, etc., to come in contact with the pacing system. Gloves will be worn whenever touching pacing electrodes to prevent transmission of microshock to patient.

  - Ensure the ends of the pacing electrode are secured into the extension cable, and that the cable is always secured into the terminals of the pulse generator.

  - Take precautions to ensure pulse generator will not fall or otherwise pull on the electrode wire or become disconnected.

  - To reduce the risk of accidental dislodgement- patient’s with transvenous wires in place may not ambulate- they must remain on bedrest or may get OOB to chair at the bedside with a provider order. Transport must be done in a bed or gurney.

  - Connect intracardiac catheters and electrode only to instruments or transducers which provide electrical isolation.

  - Wear rubber gloves when handling the catheter electrode wires.

  - Insulate exposed metal terminals on a pacemaker with a rubber glove.

  - Never touch exposed pacemaker terminals or catheter guide wire with one hand while making contact with electrically operated devices with the other hand.

  - Use plastic (non-conductive) stopcocks whenever possible on all pressure lines.

  - Evaluate ECG for presence of paced rhythm or resolution of initiating dysrhythmia with each set of vital signs and prn.

  - Monitor for pacemaker malfunctions with each set of routine vital signs and prn. See Appendix A for detailed information regarding common pacemaker malfunctions and troubleshooting.

  - Notify provider of any pacemaker problems and corrective actions taken.

INSERTION

Equipment

- Percutaneous Introducer Kit (ask provider for size); 6.0 French is routinely used

- Central line support tray

- IV NS 250 mL or 500 mL

- Sterile gloves

- 9-volt battery

- Pacemaker insertion box

- Fluoroscopy equipment and technologist if required by provider

- Lead aprons if fluoroscopy used

- Transvenous pacing wire (ask provider for type and size). If a pulmonary artery pacing catheter is being used, see Pulmonary Artery Catheter Insertion, Monitoring and Discontinuation Policy.

- Pulse generator (temporary pacemaker).

- Have crash cart and defibrillator available

Procedure

  - Insertion and Initiating Pacing

    - Ensure that defibrillator and resuscitative equipment is available at the bedside.

    - Arrange for fluoroscopy, according to provider request.  Notify Radiology of time and location of procedure.

    - Explain the procedure to the patient.

    - Verify informed consent has been obtained.

    - Perform hand hygiene and position the patient as desired by provider to expose site of insertion.

    - Ensure that each element of the CLIP tool and Universal Protocol Checklist Team Time-Out form are met.

    - Connect the patient to a bedside monitoring system and monitor ECG continuously.

    - Administer pain medication and/or sedation per provider order and Moderate Sedation policy.

    - Clip and prep area of insertion with chlorhexidine.

    - All personnel performing and assisting with procedure will wear masks, caps, goggles or face shields, sterile gowns, and gloves.

    - Insert new battery into pulse generator and check for battery signal. Record on battery the date/time of new battery insertion.

    - Assist provider with insertion of transvenous wire (pacing electrode).

    - When pacing electrode is in place, connect the distal wire to the negative terminal and the proximal wire to the positive terminal of the pulse generator. A wire connector can be used if desired.

    - Transvenous wire (pacing electrode) will be sutured in place. Apply a chlorhexidine gluconate impregnated transparent occlusive dressing per the Central Line policy. Date, time, and initial dressing.

    - Set mode, output (milliamps (mA)), sensitivity (millivolts (mV) and rate as ordered by provider.

    - If 100% capture not present, increase the mA level until 100% capture is noted. Capture is indicated by an atrial and/or ventricular response with each pacemaker impulse. To prevent fibrosis, the lowest safe mA setting should be utilized. Check pulse to ensure mechanical capture.

    - Monitor vital signs and patient symptoms until stable.

    - Observe monitor for sensing and capture. Obtain rhythm strip.

    - Notify provider of initiation of electrical pacing and the patient's status.

    - Obtain new provider orders if settings necessary for pacing are different from original order (i.e., higher mA is needed for capture than originally ordered).

  - Determining Sensitivity and Stimulation Threshold (These tests are usually done when pacing is initiated or per provider order).

  - Determining the sensitivity threshold (level at which intrinsic myocardial activity is recognized by the sensing electrodes).

    - Do not perform a sensitivity threshold test if patient has no intrinsic rhythm.

    - Perform this test with the pacer rate set at ten below the patient’s intrinsic rate and the mA at the lowest possible setting

    - Turn sensitivity down (counterclockwise, to a higher number) until sense indicator light stops flashing. This is when the pulse generator is no longer able to detect the patient’s intrinsic rhythm.

    - Slowly turn the sensitivity back up (clockwise, to a lower numeric setting) until the sense indicator light flashes with each complex and the pace indicator light stops- this is the sensing threshold.

    - Set the sensitivity dial to the number that was half the sensing threshold.

  - Determining the stimulation threshold (the amount of output required to achieve capture):

    - This test should be performed by a provider if the patient is pacemaker dependent for bradycardia.

    - Set the pacing rate 10 beats above the patient’s intrinsic rate.

    - Gradually decrease the mA from 20 until capture is lost.

    - Gradually increase mA until 1:1 capture is established. This is the stimulation or pacing threshold.

    - Set the mA at least two times above the stimulation threshold. For example, if 1:1 capture is achieved at three mA, then mA on pulse generator should be set at six mA.

DRESSING CHANGE

Change dressing according to ‘Central Venous Access Devices: Care, Maintenance, Blood Withdrawal, and Removal’ policy.

Documentation

Document the following in the patient’s medical record:

- After insertion, document the need for the pacemaker and the patient's response to the paced rhythm in the medical record.

- Chart ECG strip with date, time, pacemaker type and settings per ECG Monitoring Policy.

- Appropriateness of capture and sensing. Provider notification and corrective actions taken if capturing or sensing problems occur.

- Chart location and condition of insertion site every 12 hrs.

- If placed in an extremity, chart circulation distal to insertion site every shift and PRN.

- Chart dressing changes.

- Education on infection prevention (CLABSI Handout).

Related Policies and Procedures

- Standardized Procedure for Emergency Care and Arrhythmia Treatment

- Central Venous Access Devices: Care, Maintenance, Blood Withdrawal, and Removal

- Pulmonary Artery Catheter Insertion, Monitoring and Discontinuation

- Standardized Procedure for the Ordering of Flushing Vascular Access Device

- ECG Monitoring Policy

Appendix

- Appendix A: Pacemaker Malfunctions and Troubleshooting

Reference(s)

- Safavi-Naeini P, Saeed M. Pacemaker Troubleshooting: Common Clinical Scenarios. Texas Heart Inst J. 2016 Oct 1;43(5):415-418. doi: 10.14503/THIJ-16-5918. PMID: 27777525; PMCID: PMC5067035.

- Wiegand, Debra. AACN Procedure Manual for Critical Care 7th Edition, pp 417-429. St Louis, MO, Saunders, 2017.

| AACN’s Evidence-leveling System. Level A; Meta-analysis of multiple controlled studies or meta-synthesis of qualitative studies with results that consistently support a specific action, intervention or treatment. Level B; Well designed controlled studies, both randomized and nonrandomized, with results that consistently support a specific action, intervention, or treatment. Level C; Quality studies, descriptive or correlational studies, integrative reviews, systematic reviews, or randomized controlled trials with inconsistent results. Level D; Peer-reviewed professional organizational standards, with clinical studies to support recommendations. Level E; Theory-based evidence from expert opinion or multiple case reports. Level M; Manufactures’ recommendations only. | ##### Effective/Revision Dates for Policy # NA 10-155. Effective:; 05/02/1991; Medical Executive Committee. Revised:; 07/11/1995; Medical Executive Committee. 01/13/1998; MEC. 06/19/2002; Special Care Committee. 05/24/2005; Governing Board. 06/05/2007; Governing Board. Revised:; 06/24/2010; Governing Board. Revised:; 02/24/2011; Governing Board. Revised:; 12/05/2013; Governing Board. Revised:; 02/23/2017; Governing Board. Revised:; 01/28/2020; Policy Committee. Revised:; 03/23/2023; Quality Committee of the Governing Board (Was Transferred from SCU 10-216). Reviewed:; 06/23/2026; Policy Committee. Keywords:; Pacing Wire, Temporary Pacemaker, Insertion, transvenous, transvenous wires, pacemaker, pacer. |
| --- | --- |

Appendix A: Pacemaker Malfunctions and Troubleshooting

  - Failure to capture: Generated pacing stimulus does not initiate myocardial depolarization (pacemaker spike with no QRS or P wave after it).

- Possible causes include:

  - Lead dislodgment, low output, lead maturation (old leads), and lead or pacer failure. Epicardial wires have a higher incidence of failure after two weeks.

- Possible solutions:

  - Change patient's ECG lead to see if capture is detected in another lead (12 Lead ECG may be helpful).

  - Increase mA in increments of 1 until 100% capture achieved.

  - If problem continues, change pulse generator.

  - Failure to pace: Heart rate is below the programmed lower rate of the pacemaker and no pacer electrical output is noted on the electrocardiogram.

    - Possible causes include lead failure (lead fracture, lead displacement), pulse generator failure, battery failure, crosstalk inhibition, or oversensing.

    - Possible solutions:

- Replace battery.

- Change pulse generator.

- Double check epicardial wire to bridging cable connection.

- If the above does not resolve issue, the pacemaker may be oversensing, see below for oversensing.

  - Oversensing: Occurs when the pacemaker senses electrical signals that it should not sense, which results in inappropriate inhibition of the pacing stimulus. Oversensing resembles failure to pace/ on the monitor. If adjusting the sensitivity corrects the issue, the pacemaker is likely oversensing.

        - Possible causes: T waves, non-physiologic signals like electromagnetic interference, lead failure.

        - Possible solutions:

          - Ensure sensitivity is not set too high- turn the sensitivity down (note, when turning the sensitivity down, the number (in mV) will go up. Caution: turning sensitivity all the way down (to the highest number) will result in asynchronous pacing.

  - Undersensing: pacemaker fails to detect spontaneous myocardial depolarization, which results in asynchronous pacing. Atrial or ventricular pacing spikes arise regardless of P waves or QRS complex. Results in the appearance of too many pacing spikes or inappropriate pacing spikes (i.e., on the QRS).

  - Possible causes: Pacemaker programming problems (improper sensing threshold), insufficient myocardial voltage signal, lead, or pacer failure (fibrosis, fracture, etc.), or an electrolyte abnormality.

  - Possible solutions:

  - Ensure sensitivity is not set too low (higher number is lower sensitivity) or asynchronous pacing may occur. (The lower the number, the higher the sensitivity). Slowly turn sensitivity up (the number on the pacemaker will go down) until problem resolves. If problem continues, change battery.

  - If problem continues, change pulse generator.